## Slide 1
How to Ace
Your PT
Avoiding the common errors that fail proficiency tests.
A practical refresher for bench technicians and lab managers

## Slide 2
WHY IT MATTERS
PT is your lab's external report card
A third party sends blind samples; you analyze and report them exactly like any client samples. Your results are then compared to other lab results.
What it is
Unknown samples arrive from an accredited provider. You test them and report by a set deadline.
How it's scored
Results are compared to determine the Consensus Mean. PT Score ≥70 is Acceptable
Why Do it
Helps to monitor your laboratory quality, monitor competence and identify problems areas.
How to Ace Your PT | Avoiding Common Proficiency Testing Errors
2

## Slide 3
COMMON ERROR
Treating PT samples differently
01
WHERE LABS SLIP
Assigning your best analyst or a special instrument just for PT
Running extra replicates or “cleaning up” the method for PT
Labeling or flagging PT vials so they stand out on the bench
HOW TO ACE IT
Run PT through the same people, SOPs, and instruments as client samples
Don't over-test - PT checks your process, not a best-case run. Trust your process.
Log and store PT results in your normal workflow and records
3
How to Ace Your PT | Avoiding Common Proficiency Testing Errors

## Slide 4
COMMON ERROR
Sample handling & Contamination
02
WHERE LABS SLIP
Opening the vial before reading the kit's handling instructions
Dilution or volume errors
Trace contamination from glassware, reagents, or the bench
HOW TO ACE IT
Read the PT instructions fully before testing the sample
Double-check dilutions and units against the reporting form
Run blanks and matrix spikes to identify issues
4
How to Ace Your PT | Avoiding Common Proficiency Testing Errors

## Slide 5
COMMON ERROR
Calculation & Reporting errors
03
WHERE LABS SLIP
Wrong units on the report form
Transcription slips when copying results into the portal
Report results for the wrong sample e.g. sample 1 for sample 4
HOW TO ACE IT
Report in the exact units and format the provider requests
Have a second analyst verify every entry before submitting
Confirm entered results are for the correct sample
5
How to Ace Your PT | Avoiding Common Proficiency Testing Errors

## Slide 6
COMMON ERROR
Deadlines, Calibration & QC
04
WHERE LABS SLIP
Missing the deadline - an automatic fail, even with perfect data
Ignoring calibration drift
Ignoring out-of-control QC results
HOW TO ACE IT
Keep track of PT Reporting deadline
Verify calibration and continuing-calibration checks
Confirm blanks, spikes, and QC pass for the batch
6
How to Ace Your PT | Avoiding Common Proficiency Testing Errors

## Slide 7
BEFORE YOU SUBMIT
Your pre-submission checklist
Six quick habits that catch the errors behind most PT failures.
✓
01
Read handling instructions in the set
✓
02
Use routine methods, staff, and instruments
✓
03
Verify calibration and QC for the batch
✓
04
Check dilutions, units, and sample numbers
✓
05
Have a second person review the entries
✓
06
Submit before the deadline - then save the record
7
How to Ace Your PT | Avoiding Common Proficiency Testing Errors

## Slide 8
KEY TAKEAWAYS
Make good habits routine
1
Treat PT like a client sample
No shortcuts, no special handling — same people, methods, and instruments.
2
Most failures are avoidable
Clerical, handling, and deadline errors cause more failures than instrument limits.
3
A checklist and calendar win
A named owner, verified QC, and a second reviewer to catch problems before you submit.
Ace PT by getting it right on every sample every time.